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1. Introduction

This document describes in detail the procedures and steps used in the calibration of an F335 Scanner using the Calibration Wizard.  All scanners are calibrated and tested prior to shipping, but they may become misaligned over time and will have to be recalibrated.  Also, if a component such as the CCD board is replaced, the scanner will have to be recalibrated.
2. Tools

In order to calibrate a scanner completely, you will need a Calibration Kit (part #125096) which consists of the following tools: the Calibration Wizard software, a software tool called Maxwell, eleven calibration filmstrips, and a lens alignment tool.  The part number of each of these tools is listed in the table below.
	Part #
	Description

	125096
	Calibration Kit

	125097
	Maxwell software

	
	Calibration Wizard

	
	Bright Point strip

	
	Alignment strip

	
	Skew Target strip

	
	Film Speed strip

	
	64 Patch Color Negative

	
	64 Patch Color Positive

	
	Lens Alignment tool


The eleven calibration filmstrips may wear out over time and can be replaced by simply ordering the appropriate part number from Pakon.

3. Calibration Wizard
The Calibration Wizard software is designed to lead the operator through all the steps of calibrating an F335 scanner from beginning to end.  These steps include Film Track Test, Focus, Lens Alignment, Skew Adjustment, Tilt Adjustment, Magnification, Film Speed, Color Generation, and End-of-Line Quality Tests.
To run Calibration Wizard, first connect the scanner to the PC, start Windows 2000/XP (if not already) and turn on the scanner.  From the “Start” menu, select “Programs”, then “Pakon F-335”, then “Tools”, then “Calibration Wizard”.

One of the initial screens is shown in Figure 1.  The  Next >  button is used for stepping through each of the instructions for the current step.  The  < Back  button is used for reverting to the previous instruction.  The  Skip  button is used for skipping the current step, going to the next step.  The  Close  button will quit the program.  The other buttons are used for performing certain calibration operations.  The memo area in the lower-left of the screen displays instructions and the memo area in the lower-right displays the current status.  The button between these two fields, the “Navigation” button, can be used for going to any particular step, starting at the beginning of the step.
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Figure 1
Calibration Wizard is designed to run as a standalone program; it is not designed to run concurrently with any other software with the exception of Maxwell.  There is no need to run Maxwell separately; the Calibration Wizard will automatically start it as necessary.
4. Calibration Steps
Listed below are all the steps necessary to completely calibrate a scanner.
4.1. Initial Cleaning
Purpose:  To clean essential scanner components prior to calibration.
Tools required:  Alcohol, Q-tip, compressed air.
Preconditions:
 The scanner has been unplugged and the covers have been removed, including the internal black cover over the CCD and lens.
Steps:

· Using a Q-tip recently moistened with alcohol, gently clean the glass transfer bar under the film transport gate.
· Immediately dry with compressed air so that no alcohol remains.
· Repeat for the CCD chip.
· Repeat for the lens.
Notes:
 A new Q-tip should be used every time.  The Q-tip should not be soaked in the alcohol for more than a few seconds.
4.2. Film Track Test

Purpose:  To test and calibrate the film sense and DX sensors.
Tools required:  Bright Point strip.
Preconditions:  The scanner must be connected to the PC and turned on, and the covers have been replaced.
Steps:

· Check for message “Ready Bright Point Strip for film track test” at bottom left corner of Calibration Wizard window.
· Click  Next >  button; the scanner will start to scan.
· Insert the Pakon Bright Point Strip, emulsion side up and head first.
· When the test completes, a dialogue window will appear showing the results of the test (Figure 2).  Take note of the results.
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Figure 2
· Click   OK   in the “Track Tests Results” dialogue window.
· If Results = ‘0’ (zero), then there were no errors; continue.
· If results were nonzero, then check the “PakonFilmTrackTestLog.txt” file in the “C:\Program Files\Pakon\TLA COM Server” folder for errors.  This will indicate which sensor(s) had problems.  Replace these sensors and rerun the test.
· If no errors, click  Next >  button in the main window.
Notes:  The only way to get to this step is to (re)start Calibration Wizard; it cannot be accessed from the “Navigation” button.  IMPORTANT: If this test is run with the covers off, then it must be run again (after all the other steps), with the covers back on the scanner.
4.3. Scanner Initialization

Purpose:  To initialize the software and get the scanner nominally focused.
Tools required:  Alignment strip, Motor Speed target strip, Bright Point Strip.
Preconditions:  None.
Steps:

· Verify the Model Type, Hardware Version, and Serial Number for the scanner (Figure 3).
· If any of these is incorrect, the scanner’s EEPROM must be reset (see Notes).
[image: image4.png]Model Type:

Seral Number:
Aligrment Terget Offset
Motor Speed Target Ratio

=S =

Poducton 7

s
5]
15573

=lolx]

»

5

550

Enter Verify Scanner Identification.





Figure 3
Notes:  An F335 has model type 3350 and hardware version Production; an F335C has model type 3351 and hardware version Production.

· Enter the number written on the Alignment strip into the Alignment Target Offset box.
· Enter the number written on the Film Speed strip into the Motor Speed Target Ratio box.
· Click  Next > .
· Click the   >>   button (fast forward); the film drive starts.
· Insert the Pakon Bright Point Strip and feed into the film gate.
· Use    >    button (forward) so that the center image of the alignment strip is centered on the illuminator.
· Click on   -50   or   +50   to focus (Figure 4).  The sharper and more rectangular the bars and spikes, the better.
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Figure 4
· Click  Next > , wait for focusing to finish and follow the instructions at bottom left corner of Calibration Wizard window.
· Click  Next > ; the Windowing Adjustment screen appears.
4.4. Windowing Adjustment

Purpose:  To align the lens.
Tools required:  Bright Point Strip.
Recommended tools:  Lens Alignment tool

Preconditions:  The three outside covers have been removed, as well as the internal black metal cover, and Scanner Initialization has been completed.
Steps:

· Install the Lens Alignment tool and follow the instructions at bottom left corner of the screen (Figure 5).
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Figure 5
· Center the image between the two red bars—the light from the sprocket holes should be on the outside of both red lines.
· After initial alignment has been completed.  Movement of the lens in the up or down position may need to be completed in order to achieve an image where the left edge touches the left red bar and the same on the right side. (figure above)

· Note this lens positioning is critical for calibrating proper lens stepper settings in magnification calibration.
· Note that the left-right alignment is a CRITICAL adjustment!
· After alignment is complete, apply gentle pressure to the left and then the right side of the lens holder.  The signals should not attenuate!
· Click  Next >  when done; the lens should move down to a new position where the magnification is increased.
· Again, apply gentle pressure to the left and then the right side of the lens holder.  The signals should not attenuate!

· Click   >>   to remove the Pakon Alignment strip from the scanner.
· Click  Next > ; the Skew Adjustment screen appears.
4.5. Skew Adjustment

NOTE: This procedure is to be performed only by technicians from BTW and Pakon.  A special tool is required for setting the tension on the drive belt—this tool is only available to BTW and Pakon for manufacturing purposes.  

Purpose:  To adjust the skew of the CCD.
Tools required:  Skew Target strip.
Preconditions:  The three outside covers have been removed, as well as the internal black metal cover.
Steps:

· Disable the drive motor: loosen screws on wheels at left side of scanner behind film guide (close to center upright panel).
· Insert Pakon Skew Target.

· Slowly turn the film guide wheel – on drive motor behind film guide – to move the Skew Target strip through the film guide.
· As the Skew Target moves through the scanner, light passes through the clear areas and then is blocked by the dark bars.  

· Adjust the CCD board to reduce the skew.
· The adjustment is correct when the left and right sides of the signal rise and fall together and do not make a wave when the target is advanced.
· Figure 6 shows a poor adjustment.
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Figure 6
· Figure 7 shows a proper adjustment.
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Figure 7
· Hint: Advance the film so that the intensity rises and then drops to about halfway.  If the left side is lower, then tighten the screw on the left side of the CCD board.
· Repeat the steps above after each adjustment until result is satisfactory

· Remove the Skew Target strip from the scanner.
· Enable drive motor by retightening the film guide wheels.
· Click  Next > ; the Tilt Adjustment screen appears.
Notes:  
4.6. Tilt Adjustment

Purpose:  To adjust the tilt of the CCD.
Tools required:  Bright Point strip.
Preconditions:  The three outside covers have been removed, as well as the internal black metal cover.
Steps:

· The initial Tilt Adjustment screen is shown in Figure 8.
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Figure 8
· Click   >>   button and insert Pakon Bright Point strip, emulsion side up.

· Use    >    button to center the strip in film guide, if need be.
· Install the black metal cover to completely cover the lens and CCD board.
· Click  Next > , wait for focusing to complete.
· Remove the black metal cover.
· Follow instructions in the window for the necessary steps for adjustment.
· Repeat until adjustment is within tolerance.
· Click  Next >  when tilt adjustment is done.
· Using the “Navigation” button, redo the Windowing Adjustment.

· Reinstall the black metal cover.

Notes:  
4.7. Windowing 2 Adjustment

Purpose:  To realign the lens after skew and tilt has been removed.

Tools required:  Alignment strip.

Recommended tools:  Lens Alignment tool

Preconditions:  The three outside covers have been removed, as well as the internal black metal cover, and Scanner Initialization has been completed.

Steps:

· Install the Lens Alignment tool and follow the instructions at bottom left corner of the screen (Figure 5).
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Figure 9
· Center the image between the two red bars—the light from the sprocket holes should be on the outside of both red lines.

· Note that the left-right alignment is a CRITICAL adjustment!
· After alignment is complete, apply gentle pressure to the left and then the right side of the lens holder.  The signals should not attenuate!
· Click  Next >  when done; the lens should move down to a new position where the magnification is increased.

· Again, apply gentle pressure to the left and then the right side of the lens holder.  The signals should not attenuate!

· Click   >>   to remove the Pakon Bright Point strip from the scanner.
· Click  Next > ; the Skew Adjustment screen appears.

4.8. Magnification Calculation

Purpose:  To focus the lens precisely.
Tools required:  Alignment strip.
Preconditions:  Black metal cover is installed and nominal focus has been completed.
Steps:

· Start with the Magnification Calculation screen (Figure 10).
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Figure 10
· Insert the Pakon Alignment strip.
· Use the   >>    and    >    buttons to center the Alignment strip and achieve the results shown in Figure 10.
· Click  Next > , wait for focusing to finish.
· Click  Next > .
· Click on each of the two vertical lines in the window and drag them to the top edge of the sloping ends of the channel bar, one on the left and one on the right (see Figure 11).
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Figure 11
· Click  Next > .
· Click on   -50   or   +50   to improve the focus (the sharper, more rectangular the bars and spikes the better, Figure 12).
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Figure 12
· Click  Next > , wait for focusing to finish.
· Click on each of the two vertical lines in the window and drag them to the top edge of the sloping ends of the channel bar, one on the left and one on the right (see Figure 13).
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Figure 13
· Click  Next > , wait for focusing to finish.
· Use   >>   to remove the film strip.

· Click  Next > .
Notes:  If this test fails for any reason, make sure the stepper motors are moving and rerun the test.
4.9. Film Speed Determination

Purpose:  To adjust the motor speed.
Tools required:  Film Speed strip.
Preconditions:  Scanner Initialization has been completed and black metal cover is installed.
Steps:

· Start with the Film Speed Determination screen (Figure 14).
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Figure 14
· The calibration wizard will determine the speed for one resolution and extrapolate that speed into all resolutions.
· Insert the 35 mm Film Speed strip with the emulsion side up, head first.
· Click    >    (forward) until the edge of the last frame is at the end of the film guide, and the beginning of the black edge is hanging over the film guide.
· Click  Next > .
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· Click   >>   until the film strip is removed from the scanner.

· Click  Next > .
Notes:  The motor speed calibration will run in the order that is displayed, from left to right and top to bottom.  In the event of a failure:


a.
Remove the filmstrip from the scanner by clicking   >>  .


b.
Repeat the speed calibration steps starting with inserting the 35 mm Film Speed strip.
4.10. 64 Patch Color Matrix Generation

Purpose:  To adjust the color correction for negative and positive film.
Tools required:  64 Patch Color Negative strip and 64 Patch Color Positive strip.
Preconditions:  All covers are installed.
Steps:

· Start with the 64 Patch Color Matrix Generation screen (Figure 15); Color Matrix Negative is the first dialogue window.
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Figure 15
· Get the number at the top of the center image of the 64 Patch Color Positive strip, add “11” to the end, and enter this into the “Center Image” box.
· Enter “51112099” into the “Image Bar Code” box for color negative.
· Click  Next >  and wait for force corrections to complete.
· Look for “Ready 64 Color Patch Strip” message in the bottom left window.
· Click  Next >  and wait for scanning to start.
· Insert the 64 Patch Color Negative strip with emulsion up and head first; the film should move automatically into the scanner.

· If a “Scan Warnings” window appears, click OK, otherwise wait for transferring to complete.
· The “Verify Scan and Image Names” dialogue window appears.

· Verify images: navigate through the images by clicking the  Previous  and   Next . buttons looking for smear, inconsistent borders, and colors.

· Verify name: check that the data in the Center Image box matches the image number of the center image displaying in the window – Figure 16.
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Figure 16
· Click  Next > , wait for saving to finish and Maxwell to start (Figure 17).
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Figure 17
· Click  Production  in the Maxwell window.
· Verify image by Center Image Number and Image Bar Code; an “N/A” next to the image number indicates success.
· Wait for Maxwell to finish processing; the  Production  button should gray out when done.

· Click  Quit Production Mode .
· Click  Exit Maxwell  and wait for Maxwell to close.

· Click  Next > , remove all film, and begin Color Matrix Positive generation – Figure 18.
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Figure 18
· Click  Next >  and wait for force corrections to complete.
· Get the number at the top of the center image of the 64 Patch Color Positive strip, add “11” to the end, and enter this into the “Center Image” box.
· Enter “25511249” into the “Image Bar Code” box for color positive.

· Click  Next >  and wait for force corrections to complete.
· Look for “Ready 64 Color Patch Strip” message in bottom left window.
· Click  Next >  and wait for scanning to start.
· Insert the 64 Patch Color Positive strip with emulsion up and latent image bar codes away from the observer; film should move automatically into the scanner.

· If a “Scan Warnings” window appears, click OK, otherwise wait for transferring to complete.

· The “Verify Scan and Image Names” dialogue window appears.

· Verify images: navigate through the images by clicking the  Previous  and   Next . buttons looking for smear, inconsistent borders, and colors.

· Verify name: check that the data in the Center Image box matches the image number of the center image displaying in the window – Figure 19.
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Figure 19
· Click  Next >  and wait for saving to complete and Maxwell to start (Figure 20).

· Click  Production  in the Maxwell window.
· Verify image by Center Image Number and Image Bar Code; an “N/A” next to the image number indicates success.
· Wait for Maxwell to finish processing; the  Production  button should gray out when done.

· Click  Quit Production Mode .
· Click  Exit Maxwell  and wait for Maxwell to close.
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Figure 20
· Click Next >  .
4.11. MOF  Calibration

Purpose:  Optimize setting on the MOF cartridge scanner. This step is displayed for an F-335C. 
Tools required:  APS Film Cartridge.

Preconditions:  NONE.

Steps:

· Click  Next >  if the MOF cartridge has extracted film inserted in the scanner.  This will do a manual retract to put the film back in the cartridge.  ”Click  Skip >  if the film is completely in the MOF cartridge. (Figure 22).
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Figure 21
· The calibration will complete on its own. No other user intervention is necessary.
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· Click [image: image25.png]


 and select last menu item.
Notes:  The MOF unit has been successfully calibrated.

4.12. Scanner Matrix (Scanner Calibration Complete)

Purpose:  To finish the calibration steps.

Tools required:  None.

Preconditions:  All covers installed.

Steps:

· Click  Next > ; the “Scanner Calibration Completion” window appears (Figure 22).
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Figure 22
· Click  Close .
Notes:  The scanner has been successfully calibrated.
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